UCBM internship Weekly Work log

GitHub Repository (with full code)

Week 1 (19th Jan – 23rd Jan)

- Introduction with Professor Scotto Di Luzio and Marco along with demonstration of ALExRS working.

- Installed and configured a Windows VM using UTM on my Mac computer to run the software.

- Transferred the Release folder to the VM along with debugging multiple errors during the process of transfer

- Tested the software from the VM (All programs run and logged apart from the ALEX GUI, highlighting the need to use a Lab Windows PC)

- Observed the Mirror experiment performed by Marco and Laura and understood the state machine and transitions.

  - 5 states: reaching, grasp, hold and lift, release, turn back.

  - Release and grasp states occur when shoulder and elbow joint velocity is 0

  - Understood by experiencing what “transparent” mode means on ALEX.

- Helped with their data collection: Performed movements of all states in sequence on the exoskeleton.

- Explored the data structures (Struct_in and struct_out) documentation written by Marco for data collection.

- Translated the 90 page user manual to English and read Section 5.

- Understanding of muscle movements and corresponding joint motions

  - Since my background in physiology is limited, I spent some time understanding Osteokinetic movements along with anatomical planes and axes. (Some of my notes are below)

  - Studied Abduction, Adduction, Flexion, Extension of shoulder and elbow.

  - Mapped these movements to the joint movements of ALExRS.

Week 2 (26th Jan – 30th Jan)

- Windows machine setup:

  - Downloading the Shared memory and ALEX_GUI classes and Release folder

  - Installation of Visual Studio and required packages.

- Assisting Marco with setup and testing the 5 Bilateral tasks on ALEx:

  - Attempting to remove the end effector (Unsuccessful as the system was over compensating for gravity due to removal of mass)

  - Re-configuring the table setup to comfortably reach each target.

  - Lifting H1 and H2 task:
Measured and tested offsets (in Z direction, craniocaudal axis). Offsets in this task determine the distance moved by the end effector at which the transition from HOLD_LIFT to RELASE states occurs.

  - Movement along a plane:

Measured and tested offsets (in Y direction, anteroposterior axis). Offsets in this task determine the Y distance moved along the plane between GRASP and RELEASE states.

  - Tested the flexion and repetitive flexion cycle experiments.

  - Testing of the electrical stimulator (in progress)

- ALEx Connection and Data Collection

  - Writing code (based on sample given by Marco) for reading from the shared memory into a csv file.

Week 3 (2nd Feb – 6th Feb)

Windows machine debugging (Problem: ALExGUI_4.0 application not loading on the PC)

- Initial launch of the GUI

  - Action: Launched the GUI through Lancia_tutto.bat and directly with ALExGUI_4.0.exe, while connected to the robot and disconnected.

  - Outcome: Unsuccessful, GUI did not open and showed up as a background task in Task manager.

- Checked the system and application architecture (x64 or x86)

  - Observation: PC is x64 (64 bit), ALExGUI prefers, but does not require x86 (32 bit)

  - Action: Used corflags to force 32 bit run of ALExGUI.

  - Outcome: Unsuccessful, same as before.

- Compared Windows Settings with Marco’s PC (Because the GUI launches on his PC)

  - Action: Turned off Virtualisation based security and changed Hyper-virtualisation settings to Match Marco’s PC settings.

  - Outcome: Unsuccessful, Event viewer showed .NET and kernel errors.

- Installed all .NET sdk packages and Visual C++ redistributables on Marco’s PC

  - Action: Installed .NET SDK 10.0.101, 9.0.305, and

Microsoft VC redistributables for x64 and x86 from 2008-2014

  - Outcome: Unsuccessful, Event viewer showed SQL server errors.

- Installed SQL servers and tools present on Marco’s PC

  - Action: Installed Microsoft SQL Server Express, Native Client, LocalDB, TransactSQL ScriptDom, Support Files.

  - Outcome: Successful launch of GUI from both Lancia_tutto.bat and ALExGUI_4.0.exe!

Screenshots of setting and list of applications on Marco’s PC

(In purple: packages installed and settings that were changed on my Windows PC)

Assisting Marco with testing of Functional Electrical Stimulation

- Understood the muscles stimulated for grasping and Release and their movement.

| Name | Movement | Grasp / Release | Location |
| --- | --- | --- | --- |
| Flexor Digitorum Superficialis (FDS) | Finger Flexion | Grasp | Anterior and ulnar side Between elbow and wrist |
| Flexor Carpi Radialis (FCR) | Wrist Flexion | Grasp | Anterior and radial side Closer to elbow |
| Abductor Pollicis Brevis (APB) | Thumb opposition and abduction | Grasp | Anterior and radial side Base of thumb and near wrist |
| Extensor Pollicis Longus (EPL) | Thumb extension | Release | Posterior and ulnar side Closer to wrist |
| Extensor Digitorum Communis (EDC) | Finger Extension | Release | Posterior and central Between elbow and wrist |
| Extensor Carpi Ulnaris (ECU) | Wrist Extension | Release | Posterior and central Closer to elbow |

- Setup of EMG sensors and stimulator electrodes:

  - Assisted with locating and positioning electrodes and sensors.

  - Left hand (master): 6 EMG sensors detecting wrist and finger flexion and extension.

  - Right hand (slave): 6 channel electrical stimulation for assistance with muscle contraction.

- Measurement of MVC (Maximum Voluntary Contraction)

  - Tested each of the 6 muscles with EMG sensors on EMGWorks.

  - MVC normalised contraction is given as the current input to the stimulator.

- Measurement of minimum and maximum current

  - Measured minimum current (Imin) required for the muscle to contract.

  - Measured the maximum current (Imax) applied before discomfort.

- Tests with the 5 Bilateral tasks

  - Assisted by controlling the flow of the experiments: starting, stopping, switching between tasks and saving the data.

  - Each task was successfully performed with the stimulation assistance for grasp and release states.

  - EMG sensors measured contraction of each muscle on left hand and normalised it by MVC.

  - The normalised values mapped between Imin and Imax control the amount of current for stimulation of right hand muscles.

Week 4 (9th Feb – 13th Feb)

Data Collection from ALEx

- Collected End Effector data by performing 3 flexion cycles of the left arm (plotted below)

ALEx GUI Control Issue

- Issue: Fault Code 3011, GUI shows disconnected and I cannot control the robot from the GUI.

- Resolution attempt: Check the communication frequency (should be >20Hz)

- Observation: Frequency is 6-8 Hz (Too low to launch the GUI without faults)

- Possible solution: Create a custom GUI and test if robot can be controlled through shared memory.

Week 5 (16th Feb – 20th Feb)

Creating the Custom GUI

- Created a custom GUI for reading joint data and controlling the robot. (Using the assistance of claude.ai and reading the shared memory documentation)

- Features present in first version of GUI:

  - Displays left and right arm end effector positions and velocity

  - Displays frequency of communication

  - Joint data can be recorded in a csv file.

  - Buttons for changing the mode of the robot from startup to wearing to rehab.

- Additional features added

  - Positions and velocities of each joint and end effector are logged.

  - Joint data is recorded at a frequency of 50 Hz.

  - Freeze display that momentarily stops rapidly updating the display data to observe and note down joint positions and velocities at a time instant.

  - Record all data mode: Data logging of both all joint angles and positions, left and right.

  - Monolateral acquisition mode: (Designed for labelled training data collection)

    - Choice of arm (left or right) and logs data only from chosen arm.

    - Once data collection is started, the user performs the experiment and as the user reaches a change of state, they press the “Next State” button on the GUI.

    - In this way, the data records a label of the subtask and its start and end points.

Recording Data:

- Recorded 30 repetitions of a point to point task using ALExRS.

- Task:

  - Start from neutral position with hands on the side of the torso. (IDLE)

  - Move to a fixed point at a defined location forward and upward (GOING)

  - Return to neutral position (RETURNING)

- At each time step, the state (IDLE, GOING or RETURNING) is recorded in a csv file along with the joint data.

- The repetitions were time agnostic: the data contained variations in the duration of GOING and RETURNING states.

- The repetitions were path agnostic: the data contained variations in the trajectory traversed between the fixed points.

Week 5 (16th Feb – 20th Feb)

- Reading the papers provided: Summary and notes here

- Implementing PELT change point detection algorithm on python:

  - Detected duplicated and incorrect change points during the transition between ‘going’ and ‘returning’.

  - Reason: I stopped at the target for a few seconds and it detected that as a new state as the velocity changed from non zero to zero.

  - Solution: I removed consecutive zero values in going and returning states before applying CPD. This altered data is closer to a point to point task without the additional waiting or stationary state.

  - Observation: The performance improved

  - Inference:  New Data should be collected without movement breaks between going and returning. Additionally, the state change signifying the end of the returning state must be recorded. 
Data must have 4 states and 3 change points between them: 
START, GOING, RETURNING, STOP

PLANNED TASKS

(Based on discussions with Prof. Cordella and Prof Scotto Di Luzio)

- Connect to exoskeleton, fixed IP (Done)

- Write code for storing data in a file (Done)

- Retrieve joint data and store in file (Done)

- Connect a game using visual studio

- Store the output of the game

- Choose a task: mono-lateral, point to point, time agnostic, (Done)

- Select start and final positions of subtask in a large complex task (Done)

- Data collection of multiple repetitions of a point to point task (Done)

- Reading approaches for change point detection in Papers given by Prof. Scotto

- Implementing Few shot learning, change point detection approaches for subtask detection.

- Assistance as needed. Based on velocities and torque